# Blanchard Valley Health System      Site:  _______________________    Location: _____________________

Training / Competency Assessment for Laboratory Point of Care Testing                                 Date of Hire / Transfer:  __________________

Method:  POCT Nova StatStrip Xpress 2 Glucose Meter                                                              Date of Initial Training:  __________________

Test Analyte:   Whole blood glucose

Associate Name and Title:  __________________  ________________________ ______

(print legibly) Associate ID number:  _________________             Supervisor:   ___________________________________________________

CLIA Complexity:  Waived  (2 areas must be assessed)  (CAP approved areas 1 –  6)  (JC approved areas 1, 3, 5, 6)

___ Training assessment after initial training (new hires) and/or new method and/or new test analytes         Date: __________________

___ Competency assessment annually                                                                                                     Date Due: __________________

Trained operators using the Xpress 2 glucose meter will perform the following tasks:

- Place the test strip into the meter.  Have an understanding of expiration dates (strips:  6 months,   QC:  3 months)

- Have an understanding of how to take a blood sample from a patient following hospital protocol (capillary, arterial or venous).

(A simulated patient sample is analyzed in class.)

- Touch a drop of blood to the end of the test strip.

- Warning:  The test strip must fill completely upon touching the blood droplet.  If the test strip does not fill completely,

do not touch the test strip to the blood droplet a second time.  Discard the test strip and repeat the test with a

new test strip.

- View test data on screen.   Accept or reject the result.  Record result.  Then remove and dispose of the test strip properly.

1.____ 2.  ____ 3. ____ Perform QC testing as required.  (Properly running QC under direct observation and recording acceptable test

results meets CAP/JC competency requirements for elements 1, 2, 3.)      Record results and enter in Cerner if indicated:

Glucose              QC Level 1 Control   __________________          QC Level 3 Control   ___________________           PASS  /  FAIL    (circle)

4.____ Check batteries for corrosion and then clean and disinfect meter.  (Properly performing instrument maintenance and function checks under direct observation and documenting acceptable results meets CAP/JC competency requirements for element 4.)   Record results:

Monthly batteries check:   Acceptable?   YES / NO     (circle)

Cleaned and disinfected meter with appropriate germicidal wipe.                                    PASS  /  FAIL    (circle)

Associate Signature:  ________________________________________      Date:  _________________

Signature of Qualified Trainer/Assessor:  _________________________________    Date:  _________________

Job title of Qualified Trainer/Assessor:     ________________________________   1-5   Satisfactory / Unsatisfactory

6. ____An evaluation of problem solving skills by written exam in Net Learning with 70% passing score will be assigned to meet CAP/JC competency requirements for element 6.

Assigned: _____________    Due:  _____________   Completed:  _____________   Score: ____________                          PASS  /  FAIL    (circle)

Reviewed by:  ___________________________

Date:  ____________