Blanchard Valley Health System      Site:  _______________________    Location: _____________________

Training Checklist for Laboratory Point of Care Testing                                                              Date of Hire / Transfer:  __________________

Method:  POCT Nova StatStrip Xpress 2 Glucose Meter                                                             Date of Initial Training:  __________________

Test Analyte:   Whole blood glucose

Associate Name and Title:  _____________________________________________________

(print legibly) Associate ID number:  _________________             Supervisor:   __________________________________________________

# Describe Xpress 2 Meter

#### Highlight differences from current meter

____ 6-sec analysis time

____ 1.2 μL sample volume

____ Importance of clean and dry patient finger

____Linearity range 10-600 mg/dL

____ 2 Energizer Max AAA batteries

____ Position of strip port

Describe buttons

____ Mode

____ Arrows

____ Eject button

Describe test strips

____ Front fill

____ Gold layer

____ Do NOT under-dose or re-dose; overdosing is not possible

____ Date vials when opened; 6 month expiration after opening

____ Store at room temperature

____ Keep vial closed; affected by heat, humidity

Demonstrate how to manipulate the meter

____ Keep horizontal / keep strip well dry

____ Apply sample or QC until countdown starts

### Using the Meter

____ Introduce Quick Reference Guides

Quality

____ QC to be done each day of patient testing.  Record results on Xpress 2 QC and Maintenance log sheet.

____ Date QC vials when opened; 3 month expiration after opening

____ Store at room temperature

____ Values displayed

____ Mix vial and expel first drop

____ Keep meter on counter

____ Explain failure due to wrong vial run

____ How to follow-up on failed QC (use comments & repeat)   Process same as current practice. Comment required.  Repeat QC testing.  If necessary, obtain new controls and repeat testing.  If necessary, obtain new test strips and repeat testing.  Do not test patients until QC is acceptable.  JC/CAP accredited sites required to enter QC results and comments in Cerner.

List of BVHS Glucose Meter quality control identifiers:

BVH Cardiac Rehab:  CR - High: QC-1070391, Low: QC-1070392

BVH Laboratory:   1. OPL A (Outpatient Lab A) - High: QC-1070441, Low: QC-1070442

2. OPL B (Outpatient Lab B) - High: QC-1070451, Low: QC-1070452

Bluffton Laboratory: XLAB - High: QC-1070611, Low: QC-1070612

Carey Diagnostic Center:  CDC – High:  QC-6070002, Low:  QC-6070001

Eastern Woods Outpatient Center:  EW - High: QC-1070541, Low: QC-1070542

Ottawa Diagnostic Center:  ODC - High: QC-2070621, Low: QC-2070622

Wound Care Solutions:   1. WC A - High: QC-1070531, Low: QC-1070532

2. WC B - High: QC-1070681, Low: QC-1070682

Run a sample and report results

____ Check strip lot for expiration date.

____ Comment (Provider notified) required for critical result.

____ How to test an isolation patient* Current practice unchanged.  Trained operators should only take glucose      meter and testing strip in room.  OK to put glucose meter in clear zip lock bag with only testing strip exposed.  However, glucose meter must still be disinfected.

Review / Look up Results

____ Memory 400 samples

Batteries Inspection

____ Inspect batteries monthly for expiration date for swelling, cracking, leaking, or any white powder on the outside of the batteries or in the battery case.  Record findings on Xpress 2 QC and Maintenance log sheet.  Replace batteries if needed and document on Xpress 2 QC and Maintenance log sheet.

____ Update date and time

Meter Cleaning and Disinfecting

____ Clorox Healthcare Bleach Germicidal Wipes

____ Surface Contact Time:  Ensure the meter surface stays wet Clorox Wipe: 1 minute

Exception: Known diagnosis of C. diff infection 3 minutes

Allow to air dry for an additional 1 minute

Additional Agenda Items for Super User Training:

____ Cleaning and avoiding the strip port

____ Troubleshooting failed Quality Control

____ Other Error Code troubleshooting

____ Glucose test strip technology

____ Question and answer session

____ Distribute “StatStrip Xpress 2 Meter Training” Printouts (for Super Users to use as a reference when training)

Associate Signature:  _________________________________________________________      Date:  ______________

Signature of Qualified Trainer/Assessor:  __________________________________________     Date:  ______________

Job title of Qualified Trainer/Assessor:     __________________________________________