Blanchard Valley Health System: Point of Care Testing

HemoCue 801 Procedure

PRINCIPLE:

The HemoCue Hb 801 System consists of an analyzer together with microcuvettes. The microcuvette serves both as a pipette and as a sample carrier. Blood is drawn into the microcuvette

cavity by capillary action. The measurement takes place in the analyzer, which measures the

absorbance of whole blood at an Hb/HbO2 isosbestic point (506 nm), and at a wavelength

(880 nm) to compensate for possible interfering background (e.g. turbidity).

The HemoCue Hb 801 System is calibrated against the hemiglobincyanide (HiCN) method, the

international reference method recommended by ICSH (International Council for

Standardization in Haematology), for the determination of the hemoglobin concentration in

blood.

CLINICAL SIGNIFICANCE:

The test is for quantitative determination of hemoglobin in capillary or venous whole blood (K2EDTA and Li-Heparin) in point-of-care settings. The HemoCue® Hb 801 System is intended to be used to determine the hemoglobin concentration for adults, adolescents, children, and infants above 1 month old. The HemoCue® Hb 801 System is for professional in vitro diagnostic use only

INDICATIONS FOR TESTING:

- Shortness of breath without identifiable cause

- Lightheadedness, dizziness, syncope

  - Should be ordered in conjunction with orthostatics, EKG, POC glucose if diabetic or concern for hyper/hypoglycemia

- Fatigue: acutely for days to weeks

- Melena/Hematochezia

- If concern for GI bleed should perform Hemoccult additionally

- If symptomatic or if vitals abnormal (tachycardic, hypotensive, tachypneic)

- Palpitations, tachycardia, new onset arrhythmias

- Pregnancy with new onset cardiovascular symptoms

  - If HDS and non-critical result, follow up with OB/GYN

  - If HD unstable or critical result, escalate care to ER

INDICATIONS FOR NOT TESTING:

- Monitoring for anemia of chronic disease

  - Advise patient to follow up with PCP or specialist who is responsible for monitoring

  - Refer PCP if not established

- Chronic fatigue

  - Follow up PCP or refer to PCP if not established for full lab work

- Routine screening for infants/children

  - Follow up Pediatrician or refer if not established

- Patient requests bloodwork as a basic check

  - Follow up PCP or refer if not established

SPECIMEN:

Capillary or venous whole blood may be used. K2EDTA and Li-Heparin should be used as the anticoagulant (for venous samples), preferably in solid form to avoid dilutional effects. If refrigerated, allow the sample to reach operating conditions. Mix all specimen tubes thoroughly on a mechanical mixer for at least 2 minutes or invert the tube 8-10 times by hand.

EQUIPMENT AND MATERIALS:

Equipment:  HemoCue Hb 801 system

Materials:

- HemoCue® Hb 801 Microcuvettes

- Eurotrol® Hb 301/801 HemoTrol Duo

CALIBRATION:

The system is factory calibrated and needs no further calibration.

QUALITY CONTROL:

Internal:  The HemoCue Hb 801 Analyzer has an internal quality control, a self-test. It automatically

verifies the performance of the analyzer every time the analyzer is turned on, when the microcuvette holder is put back into place after removal, and every hour when in use

External: Eurotrol® Hb 301/801 HemoTrol Duo controls should be run with each new lot or shipment of   microcuvettes or every 30 days.  The ranges for the QC are printed on the package insert for the Eurotrol controls. These package inserts should be saved and attached to the quality control record. Run each level of QC, record it on the QC log, and check to see that the result falls within the designated range printed on the package insert. The control ranges will change with each new lot of quality control.  Please check that the control lot matches the range printed on your form. Controls expire 30 days after opening. Controls may be stored at room temperature after opening.

If one or both levels of QC fails, repeat the failed QC. If the QC passes, patient samples may be analyzed. If the control is still out of range, open a new vial of control.  If the control continues to be out of range, open a new canister of microcuvettes. If the control is still out of range, contact BVH Point of Care Coordinator. Call 419.423.5247.

PROCEDURE-STEPWISE:

Fingerstick procedure:

- The analyzer should be in the “ready state” prior to filling the microcuvette.

- Remove a microcuvette from the vial and recap the vial immediately.

- Make sure the patient’s hand is warm and relaxed. Use only the middle or ring finger for sampling. Avoid fingers with rings on. Sample at the side of the fingertip for best blood flow and comfort.

- Clean the puncture site and allow to dry.

- Using your thumb, lightly press the finger from the top of the knuckle towards the fingertip to stimulate blood flow.

- While applying light pressure towards the fingertip, puncture the fingertip using a high flow lancet.

- Using a dry gauze or other lint-free tissue, wipe away the first two or three large drops of blood, applying light pressure as needed again until another drop of blood appears.

- Make sure that the drop of blood is big enough to fill the microcuvette completely. Hold the microcuvette opposite the filling end and introduce the microcuvette tip and fill the microcuvette in one single step. Do not refill.

- Wipe off any excess blood from the outside of the microcuvette using a clean, lint-free wipe. Do not touch the open end of the microcuvette.

- Look for air bubbles in the filled microcuvette. If present, discard the microcuvette and fill a new microcuvette from a new drop of blood.

- Place the microcuvette into the cuvette holder and start the measurement as soon as possible, but no later than 40 seconds after filling the microcuvette, by gently pressing the cuvette into the microcuvette holder.

- The hemoglobin value is displayed within a second. The result will remain on the display for 10 seconds after the measurement.

- Remove and discard the microcuvette in an appropriate biohazard container, following local procedures for disposal.

- When the display shows the “ready state” symbol, the analyzer is ready for the next measurement.

- Always handle blood specimens with care, as they might be infectious.

- See the Operating Manual for information on Repeat Testing. For further information, contact HemoCue.

Venipuncture and Quality Control procedure:

- For venous samples, use anticoagulant (K2-EDTA or Li-Heparin). Venous samples can be stored

and transported for up to 24 hours, at room temperature or in refrigerator, before measuring.

- If refrigerated, allow the sample to reach operating conditions. Mix thoroughly on mixer for at least 2 minutes, or invert manually 8–10 times.

When mixing controls, allow them to warm room temp for 15 minutes if refrigerated.  Directly before use, gently mix by hand for at least 30 seconds to completely mix the contents.

- Dispense a  drop of blood onto a  hydrophobic surface (for controls, discard the first drop of QC material), e.g. plastic film like parafilm, using a suitable transfer device.

- Fill the microcuvette in one continuous process by capillary action. Inspect for bubbles in microcuvette, if found restart with new device. Wipe off excess blood using a lint free cloth without touching the open end of the microcuvette.

- Place microcuvette into the holder within 40 seconds of filling the microcuvette.

- Make sure the analyzer is in Ready State. Insert the filled microcuvette into the microcuvette holder and press down. The result will be displayed within a second.

- When a result is displayed, remove and discard the microcuvette.

REPORTING RESULTS:

Normal Values:

| Hemoglobin |  | 0-1 Day | 19.0 - 30.0 | g/dL |
| --- | --- | --- | --- | --- |
|  |  | 1-3 Days | 14.5 - 22.5 |  |
|  |  | 4 Days - 1 Month | 13.4 - 19.8 |  |
|  |  | 1-2 Months | 10.7 - 17.1 | g/dL |
|  |  | 2-6 Months | 9.4 - 13.0 |  |
|  |  | 6 Months-2 Years | 10.5 - 13.5 |  |
|  |  | 2-6 Years | 11.5 - 13.5 |  |
|  |  | 6-12 Years | 11.5 - 15.5 |  |
|  | Male | 12-18 Years | 13.0 -16.0 | g/dL |
|  | Female | 12-18 Years | 12.0 - 16.0 |  |
|  | Male | 18 Years and Older | 13.5 - 17.5 |  |
|  | Female | 18 Years and Older | 12.0 - 16.0 |  |

Critical Hemoglobin is <6.6g/dL and >19.9 aged 12 and over

<6.9g/dL and > 20.8  1 month to 12 years

<9.5g/d/L and > 22.5  0 Day- 1 month

Critical Hgb will be reported to the provider in a timely manner and documented appropriately in the chart with time notified and the Hgb result. EMS should be contacted as method of transport due to risk of hemodynamic instability, ACS or CVA.

PROCEDURAL NOTES:

- Microcuvettes are to be stored at 50-104°F (10-40°C). The vial should be kept tightly capped and microcuvettes should be removed as needed for testing just prior to use. Microcuvettes in the vial (opened or unopened) are stable until the expiration date printed on the package. Microcuvettes in the vial (opened or unopened) can be stored for a shorter period of time (6 weeks) between 0-122°F (-18-50°C).

- The operating temperature of the HemoCue® Hb 801 Analyzer is 50-104°F (10-40°C)

- The measuring range of the analyzer is 1.0 – 25.6 g/dL. Results will display as LLL (when below) and HHH (when above) the measuring range.

- Refer to Troubleshooting (Section 15) of the Operating Manual for additional details.

LIMITATIONS OF THE PROCEDURE:

- The measurement needs to be started no later than 40 seconds after filling the microcuvette.

- Venous samples (K2EDTA or Li-Heparin) can be stored and transported for up to 24 hours in room temperature or refrigerator, before measuring.

- Refer to Limitations (Section 18) of the Operating Manual for additional details, including interfering concentrations.

REFERENCES:

HemoCue® Hb 801 Operating Manual (901912 190204 US)

HemoCue® Hb 801 System Package Insert (151904 191112 US)

Medical Director Approval

Approved by __________________________  Date  _____________________