Training Checklist for Laboratory Point of Care Testing       Date of Initial Training:______________

Method:  POCT HemoCue 801		 Test Analyte:   Hemoglobin

Associate Name and Title:  ___________________________________ ______   (print legibly)

Associate ID number:  _______________

I. Sample Collection and Preparation

_____1.Capillary or venous whole blood may be used. K2EDTA and Li-Heparin should be used as the 	anticoagulant (for venous samples), preferably in solid form to avoid dilutional effects. If 	refrigerated, allow the sample to reach operating conditions. Mix all specimen tubes thoroughly 	on a mechanical mixer for at least 2 minutes or invert the tube 8-10 times by hand.

II. Reagent Storage and Stability

_____1. Cuvettes are stored at room temperature until manufacturer's expiration date.

_____2. Unopened QC material is stored in the refrigerator until manufacturer's expiration date, once 	opened the expiration date is 30 days.

III. Safety including Infection Control / Prevention

_____1.  Wear gloves during testing events.

_____2.  Clean hands with effective antimicrobial agent.

_____3.  Discard materials designed for single use in appropriate container.

IV. Quality Control

_____1. Frequency of performing Quality Control: 2 Levels with new lot, new shipment, every 30 days 	and with each new untrained operator.

VI. TESTING: Patient and QC

_____1. Pull cuvette holder out to the loading position and turn on the analyzer.

_____2. Puncture the finger, wipe away the first 2-3 drops. Apply light pressure until a large enough        	    	sample is available to fill the cuvette in a continuous process.

_____3.Wipe off any excess blood from the outside of the microcuvette, in a “butter knife” fashion, look 	for air bubbles, if present, discard.

_____4. Place the microcuvette into the cuvette holder and start the measurement within 40 seconds

_____5. The hemoglobin value is displayed within a second

VI. Recording and Reporting results

_____1.  Hemoglobin critical test results:

- <6.6g/dL and >19.9 aged 12 and over

- <6.9g/dL and > 20.8  1 month to 12 years

- <9.5g/d/L and > 22.5  0 Day- 1 month

VII. Cleaning and Disinfecting

_____1. Weekly, clean the cover glass with a swab moistened with water.

_____2. Know how to clean the outside of the analyzer

Associate Signature:  _______________________________________________      Date:___________

Signature of Qualified Trainer/Assessor:  ________________________________     Date:____________

Job title of Qualified Trainer/Assessor:     ________________________________